Laboratory: _________________________________________ Date Implemented___________

Address: ___________________________________________

___________________________________________________

___________________________________________________

ProAdvantage® by NDC Urine/Serum Pregnancy Cassette Device

- Test Principle

The Urine/Serum Pregnancy Cassette Device is a rapid chromatographic immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in urine or serum to aid in the early detection of pregnancy. The test utilizes a combination of antibodies including mouse monoclonal anti-hCG antibodies and goat polyclonal anti-hCG antibodies to selectively detect elevated levels of hCG. The assay is conducted by adding urine or serum specimen to the specimen well of the test device and observing the formation of colored lines. The specimen migrates via capillary action along the membrane to react with the colored conjugate.

Positive specimens react with the specific colored antibody conjugates and form a colored line at the test line region of the membrane. Absence of this colored line suggests a negative result. To serve as a procedural control, a colored line will always appear at the control line region if the test has been performed properly. S

- Specimen Collection/Treatment

| A. Specimen: | Acceptable: Urine and serum specimens Unacceptable: Specimens from other sources |
| --- | --- |
| B. Specimen Collection: | Urine: Clean, dry collection container. A first morning specimen is preferred; however urines collected any time of day may be used. Serum: Collect aseptically into a clean tube without anticoagulants. |
| C. Specimen Storage: | Urine and serum samples may be stored 2-8°C up to 48 hours. For prolonged storage, specimens may be frozen and stored below -20°C. Frozen specimens should be thawed and mixed before testing. Urine: Specimens exhibiting visible precipitates should be centrifuged, filtered, or allowed to settle to obtain a clear specimen for testing. Serum: Separate the serum from blood as soon as possible to avoid hemolysis. Use clear non-hemolyzed specimens when possible. |
| D. Handling Precautions: | All specimens should be considered potentially hazardous and handled in the same manner as an infectious agent. The test cassette should be discarded in a proper biohazard container after testing. If samples are to be shipped, they should be packed in compliance with federal regulations covering the transportation of etiologic agents. |

- Reagents and Equipment

- Reagents and Materials Provided

| - Component | - Content | - Quantity |
| --- | --- | --- |
| - Test Devices | - The test device contains anti-hCG particles and anti-hCG coated on the membrane | - 40 |
| - Disposable specimen dropper |  | - 40 |
| - Package insert |  | - 1 |

- Materials not Provided

- Specimen collection container

- Timer

- Storage and Stability

Store as packaged at 2-30°C.  The test device is stable through the expiration date printed on the sealed pouch or canister label.  The test device must remain in the sealed pouch or closed canister until use.  DO NOT FREEZE.  Do not use beyond the expiration date.  Note:  Once canister has been opened, the remaining test devices are stable for 12 months.

- Quality Control

Internal procedural controls are included in the test.  A red line appearing in the control region (C) is the internal procedural control.  It confirms sufficient specimen volume and correct procedural technique.  A clear background is an internal negative background control.  If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test.

It is recommended that a positive hCG control (containing >25 mIU/mL hCG) and a negative hCG control (containing “0” mIU/mL hCG) be evaluated to verify proper test performance with each new lot, each new shipment, monthly as a check on storage, each new untrained operator and as otherwise required by your lab’s internal quality system procedures.

- Precautions

- For professional in vitro diagnostic use only.  Do not use after the expiration date.

- The test device should remain in the sealed pouch or closed canister after use.

- All specimens should be considered potentially hazardous and handled in the same manner as an infectious agent.

- The test device should be discarded in a proper biohazard container after testing.

- Test Procedure

Allow the test device, urine or serum specimen and/or controls to equilibrate to room temperature (15-30°C) prior to testing.

- Remove the test device from the sealed pouch and use it as soon as possible.

- Place the test device on a clean and level surface. Hold the dropper vertically and transfer 3 full drops of urine or serum (approx. 100 μL) to the specimen well (S) of the test device, and then start the timer. Avoid trapping air bubbles in the specimen well (S). See the illustration below.

- Wait for the red line(s) to appear. Read the result at 3 minutes when testing a urine specimen, or at 5 minutes when testing a serum specimen. It is important that the background is clear before the result is read.

- Interpretation of Test Results

POSITIVE: Two distinct red lines appear.  One line should be in the control region (C) and another line should be in the test region (T).

NEGATIVE: One red line appears in the control region (C).  No apparent red or pink line appears in the test region (T).

INVALID: Control line fails to appear.  Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure.  Review the procedure and repeat the test with a new test device.  If the problem persists, discontinue using the test kit immediately and contact your local distributor.

NOTE:  The intensity of the red color in the test line region (T) will vary depending on the concentration of hCG present in the specimen.  However, neither the quantitative value nor the rate of increase in hCG can be determined by this qualitative test.

- Limitations

- Very dilute urine specimens, as indicated by a low specific gravity, may not contain representative levels of hCG. If pregnancy is still suspected, a first morning urine specimen should be collected 48 hours later and tested.

- False negative results may occur when the levels of hCG are below the sensitivity level of the test. When pregnancy is still suspected, a first morning urine or serum specimen should be collected 48 hours later and tested.

- Very low levels of hCG (less than 50 mIU/mL) are present in urine and serum specimen shortly after implantation. However, because a significant number of first trimester pregnancies terminate for natural reasons,5 a test result that is weakly positive should be confirmed by retesting with a first morning urine or serum specimen collected 48 hours later.

- A number of conditions other than pregnancy, including trophoblastic disease and certain non-trophoblastic neoplasms including testicular tumors, prostate cancer, breast cancer, and lung cancer, cause elevated levels of hCG.6-7 Therefore, the presence of hCG in urine or serum specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.

- As with any assay employing mouse antibodies, the possibility exists for interference by human anti-mouse antibodies (HAMA) in the specimen. Specimens from patients who have received preparations of monoclonal antibodies for diagnosis or therapy may contain HAMA. Such specimens may cause false positive or false negative results.

- This test provides a presumptive diagnosis for pregnancy. A confirmed pregnancy diagnosis should only be made by a physician after all clinical and laboratory findings have been evaluated.

- Expected Values

Negative results are expected in healthy non-pregnant women and healthy men. Healthy pregnant women have hCG present in their urine and serum specimens.  The amount of hCG will vary greatly with gestational age and between individuals.

The Urine/Serum Pregnancy Cassette Device has a sensitivity of 25 mIU/mL, and is capable of detecting pregnancy as early as 1 day after the first missed menses.

- Performance Characteristics

  - Accuracy

A multi-center clinical evaluation was conducted comparing the results obtained using the Urine/ Serum Pregnancy Cassette Device and another commercially available serum/urine membrane hCG test. The urine study included 159 specimens and both assays identified 88 negative and 71 positive results. The serum study included 73 specimens and both assays identified 51 negative and 21 positive and 1 inconclusive results. The results demonstrated a 100% overall agreement (for an accuracy of >99%) of the Urine/ Serum Pregnancy Cassette Device when compared to the other urine/ serum membrane hCG test.

  - Sensitivity and Specificity

The Urine/Serum Pregnancy Cassette Device detects hCG at a concentration of 25 mIU/mL or greater. The test has been standardized to the W.H.O. Third International Standard.  The addition of LH (300 mIU/mL), FSH (1,000 mIU/mL), and TSH (1,000 mIU/mL) and positive (25 mIU/mL hCG) specimens showed no cross-reactivity.

  - Interfering Substances

The following potentially interfering substances were added to the hCG negative and positive specimens.  All substances listed in mg/dL unless otherwise noted.

| Acetaminophen | 20 | Ethanol | 1% |
| --- | --- | --- | --- |
| Acetone | 1,000 | Estriol | 2 |
| Acetylsalicylic Acid | 20 | Estrone 3-Sulfate | 10 |
| Acetoacetic Acid | 2,000 | Gentisic Acid | 20 |
| Ampicillin | 20 | Glucose | 2,000 |
| Ascorbic Acid | 20 | Hemoglobin | 1,000 |
| Atropine | 20 | Heroin | 1 |
| Albumin | 2,000 | Ibuprofen | 20 |
| β-Hydroxybutyrate salt | 2,000 | Methadone | 10 |
| Benzoylecgonine | 10 | Methamphetamine | 10 |
| Bilirubin | 20 | Methanol | 10% |
| Brompheniramine | 20 | Morphine | 0.6 |
| Caffeine | 20 | Oxalic Acid | 40 |
| Cannabinol | 10 | Phenothiazine | 20 |
| Chlomiphene | 100 | Phenylpropanolamine | 20 |
| Cocaine | 10 | Pregnanediol | 2 |
| Codeine | 10 | Salicylic Acid | 20 |
| Cholesterol | 500 | Tetracycline | 20 |
| Creatine | 20 | Triglycerides | 1,200 |
| Dextromethorphan | 20 | Theophylline | 20 |
| DMSO | 5% | Urea | 2,000 |
| EDTA | 80 | Uric Acid | 20 |
| Ephedrine | 20 |  |  |

None of the substances at the concentration tested interfered in the assay.

- References

- Batzer FR. “Hormonal evaluation of early pregnancy”, Fertil. Steril. 1980; 34(1): 1-13

- Catt KJ, ML Dufau, JL Vaitukaitis “Appearance of hCG in pregnancy plasma following the initiation of implantation of the blastocyte”, J. Clin. Endocrinol. Metab. 1975; 40(3): 537-540

- Braunstein GD, J Rasor, H. Danzer, D Adler, ME Wade “Serum human chorionic gonadotropin levels throughout normal pregnancy”, Am. J. Obstet. Gynecol. 1976; 126(6): 678-681

- Lenton EA, LM Neal, R Sulaiman “Plasma concentration of human chorionic gonadotropin from the time of implantation until the second week of pregnancy”, Fertil. Steril. 1982; 37(6): 773-778

- Steier JA, P Bergsjo, OL Myking “Human chorionic gonadotropin in maternal plasma after induced abortion, spontaneous abortion and removed ectopic pregnancy”, Obstet. Gynecol. 1984; 64(3): 391-394

- Dawood MY, BB Saxena, R Landesman “Human chorionic gonadotropin and its subunits in hydatidiform mole and choriocarcinoma”, Obstet. Gynecol. 1977; 50(2): 172-181

- Braunstein GD, JL Vaitukaitis, PP Carbone, GT Ross “Ectopic production of human chorionic gonadotropin by neoplasms”, Ann. Intern Med. 1973; 78(1): 39-45

- ProAdvantage® by NDC Urine/Serum Pregnancy Cassette Device Package Insert

Test Procedure Review

| Supervisor | Date Reviewed | Supervisor | Date Reviewed |
| --- | --- | --- | --- |